## Sheet1
Audit Observation report
| Auditor's Nmae |  | Mr. Sudhakar D & Shilpa Pittale
| Plant |  | Daman
| Auditee Nmae |  | Mr. Indresh Kumar, Mr. Hitesh patel, Mr. Puranmal sharma, Mr. Sharad dahake, Mr. Mrudang & Mr. Bhavesh rajyaguru
| Audit date |  | 10/04/2017 to 13/04/2017
Sl. No. | Department | Area | Observations | Resposibilities | Action Plan | Date of Completion | Remarks
Quality
1 |  | Sampling booth 1.1 | Weighing balance calibration validity found over due | Nilesh /Mahesh | Report verified and confirmed Calibration was perfomed on time but status label were not replaced. However status label has been replaced | Closed on 10/04/2017
|  | 1.2 | RLAF IQ, OQ & PQ document is not avaialble | Indresh / Hitesh | Retrospective Protocol & Report of IQ & OQ shall be prepared and performed as per protocol | 29/05/2017
2 |  | SMF 2.1 | SMF to be revised & update with current system and avaialbe manpower/external laboratory /JD/Nature of product manufactured | Indresh / Hitesh | SMF Shall be revised and will be updated | 24/04/2017
|  | 2.2 | Product list to be updated | Indresh / Hitesh
|  | 2.3 | Contract workmen shall be removed from SMF | Indresh / Hitesh
|  | 2.4 | Overall SMF to be reviewed and updated | Indresh / Hitesh
3 |  | Quality mannual 3.1 | Quality Mannual - To be reviewed and update as per current parctices | Indresh / Hitesh | Quality Manual Shall be revised and will be updated | 24/04/2017
4 |  | VMP 4.1 | VMP- Schedule shall be make in SOP format with inclusion of status | Indresh / Hitesh | VMP Schedule shall be revised with inclusion of formating number | 24/04/2017
5 |  | BMR 5.1 | BMR : VR06- Issue sign/date record shall be maintained | Indresh / Hitesh | Training will be porivded and same will be maintained accordingly | 22/04/2017
|  | 5.2 | Next review shall be removed from BMR | Indresh / Hitesh | BMR Shall be revised by intiating CCF to remove the next review date | 05/05/2017
|  | 5.3 | Controlled copy seal should put with sign &date BMR: As per BMR mentioned for water qty as Q.S. , it is not defined actual qty. weight in gram/litre | Indresh / Hitesh | BMR Shall be revised with inclusion of actual water qty. | 05/05/2017
|  | 5.4 | Review sign/date is not avaialbe some of pages in BMR | Indresh / Hitesh | Training shall be given to concern person for GDP | 06/05/2017
|  | 5.5 | Specimen seal (Controlled /Master) shall be defined in the SOP for location which corner suppose to be putted | Indresh / Hitesh | Training shall be given to concern person for GDP | 06/05/2017
|  | 5.6 | BMR : Revison - 06- CCF is not avaialbe | Indresh / Hitesh | CCF shall be raised and will be updated | 03/05/2017
|  | 5.7 | PM- Standard qty. to be defined in BMR | Indresh / Hitesh | BMR Shall be revised and will be updated | 05/05/2017
6 |  | CCF 6.1 | CCF Format : Persent status, proposed status , Impact analysis to be included in fromat SOP to be revised | Indresh / Hitesh | SOP QA-053/I shall be revised with inclusion of persnet status , proposed change and product impact analysis | 06/05/2017
|  | 6.2 | CCF - Some of CCF documents does not have QA manager sign & Princial company sign | Indresh / Hitesh | CCF shall be updated with sign/date | 29/04/2017
7 |  | Raw material 7.1 | PCMX: GRN no. is not avaialbe The analyst has reported test comply for solubility test but Analyst has confirmed not perfomed for fixed oil parameters Sample qty as per SOP shall be sampled 100 g but actual sample collected only 50 g | Indresh / Hitesh | Training shall be given to concern person for sampling and testing of raw materials as per SOP | 22/04/2017
|  | 7.2 | RB Specifications: Spec received through mail but it is not verifed by QA | Indresh / Hitesh | Received Specification shall be verifed by QA | 10/05/2017
|  | 7.3 | IR Test- Dodel laboratory performed the test and given the report comply as per standard but standard chromatogram is not shared | Indresh / Hitesh | External Lab IR report shall be received with Standard chromatogram | 24/04/2017
8 |  | Process validation 8.1 | PV- Protocol is not apporved by QA , Principal company & Production manager | Indresh / Hitesh | Protocol shall be signed and upadated | 22/04/2017
9 |  | Cleaning /Sanitization 9.1 | Swab test - Swab test being perfomed as per frequency and micro results comes after 5 days However During analysis results is not awaited and production is contineuing | Indresh / Hitesh | Conclusion of Equipment cleaning validation report shall be revised | 05/05/2017
10 |  | SOP/Training 10.1 | SOP- PR-01,PR-04, PR-05 and PR-06 These SOP'S are revised based on CCF and training record is not avaialbe Document control : Current and obsolete version found as master. Old version suppose to be made obsolete | Sharad /Hitesh | Training shall be given to all concern person about GDP | 29/04/2017
11 |  | Finished Product Analysis Protocol for C&C FFW 11.1 | Specific Gravity shall be recorded after Fill volume calculation | Indresh / Hitesh | Finished product analysis Protocol format shall be revised | 15/05/2017
12 |  | Approved Vendor List of J&J 12.1 | Distributor list is not updated with changes | Indresh / Hitesh | Approved vendor list shall be revised and will be updated | 28/04/2017
13 |  | Rinse Water Requisition Slip13.1 | Analysis By dated 29/03/2017 and Approved Dated 19/03/2017 | Indresh / Hitesh | It was typo error and training has been given for GDP and same document is corrected. | Closed on 14/04/2017
14 |  | Dettol Antiseptic Liquid BMR 14.1 | Actual weight of material shall not predifned in the BMR | Indresh / Hitesh | BMR Shall be revised with actual weight and blank space will be filled manually | 05/05/2017
16 |  | HVAC Qualification 16.1 | Location of filter is not difined in HVAC Qualification report(Autocal). | Indresh / Hitesh | Master document for Filter ID of HVAC and location shall be prepared. | 2017-10-05
17 |  | Master of Master 17.1 | Master of Master is not available | Indresh / Hitesh | Master of Master shall be prepared | 15/05/2017
18 |  | Complaint Investigation Report 18.1 | Pack size, Exp. Date is not recorded in Log book as per SOP | Indresh / Hitesh | Complaint Log book shall be updated as per SOP | 28/04/2017
|  | 18.2 | Category of complaint recorded is Major in investiogation report but as per SOP it's an Critical complaint | Indresh / Hitesh | Investigation report shall be corrected. | 21/04/2017
19 |  |  |  | Indresh / Hitesh
19 |  | Failure Investigation Procedure 19.1 | Responsibility: Quality Assurance manger is responsible for informing to the Director of VVF (India) Ltd. Daman | Indresh / Hitesh | SOP QA-069/C shall be revised and responsibility shall be included as per current activity. | 25/04/2017
|  | 19.2 | Quality Assurance Head of Principal Company is responsible for investigation | Indresh / Hitesh
HR
1 |  | JD 1.1 | JD - JD is not avaialbe by name based | Mrudang /Ketan | Prepare New SOP to incorporate the same | 24/04/2017
2 |  | Organogram 2.1 | Organogram is not avaialbe in SOP format | Mrudang /Ketan | Prepare New SOP to incorporate the same | 24/04/2017
3 |  | Specimen 3.1 | Specimen signature is not avaialble | Mrudang /Ketan | Prepare New SOP to incorporate the same | 24/04/2017
4 |  | Agreement 4.1 | Agreemment is not avaialbe for medical checkup | Mrudang /Ketan | No agreement done in past so for this matter we will mentioned this points in existing SOP with change control formalities | 27/04/2017
5 |  | Health checkup report 5.1 | Health cheup report- Conclusion is not mentioned | Mrudang /Ketan | Already send to party back and they will included this line in all reports and send us back with their sign and stamp | 02/05/2017
Production
|  | 1 | To be change the use of hand sanitizer with 70% IPA | Mr. Sharad | Raise change control and revise the SOP & documents Raise IPA purchase requestion & start use of IPA. | 18/04/2017
|  | 2 | AHU 1 is common for manufacturing and change room | Mr. Sharad | Check IQ/OP/PQ | 13/04/2017
|  | 3 | Convert room name - RM entry to day stock store area | Mr. Malay | Make tag as " STORE - DAY STOCK" and display in room | 13/04/2017
|  | 4 | Calibration of Standard weights / filling machine / weighing balance to be add in calibration schedule | Mr. Malay | add in calibration schedule | 16/04/2017
|  | 5 | display label not found for - STANDARD WEIGHTS | Mr. Malay | Make tag and display "STANDARD WEIGHTS" ↓ | 2017-12-04
|  | 6 | RM entry door not in locked condition | Mr. Malay | Keep door in locked condition when not in use | 2017-12-04
|  | 7 | RM drums and bags not seems proper clean | Mr. Malay | assure the cleaning of RM drums / bags before taking in STORE - DAY STOCK | 2017-12-04
|  | 8 | drum lorry - not proper cleaned & painted | Mr. Sharad / Jignesh | smooth the surface & paint the lorry | 16/04/2017
|  | 9 | RM Qty. issued by store is not match with reqd. Qty. mention in BOM. Also no tag for left over Qty. balance in loose drum / bag | Mr. Sharad | Assure - store will issue required qty.as per BOM, in presence of IPQA & production and left over qty. will keep in Store - day stock area with balalnce qty. tag | 2017-12-04
|  | 10 | Unused hole available in dispensing pipe rack | Mr. Malay / Mr. Jignesh | Plug the unused hole of dispensing pipe rack | 16/04/2017
|  | 11 | old record / loose paper / tags / knife found in table (dispensing area) | Mr. Sharad / Mr. Malay | 5S for dispensing table | 15/04/2017
|  | 12 | MB alert switch in dispensing booth- function details to be provide | Mr. Sharad | contact to OEM & try to get details | 18/04/2017
|  | 13 | As per SOP of RM dispensing - all material weighed under dispensing booth but intact drums / bags are weighed on another weighing balance (capcity 600kg) | Mr. Sharad | Raise change control and revise the SOP | 18/04/2017
|  | 14 | delay in weekly cleaning of filters of dispensing booth | Mr. Sharad | Raise change control and make the SOP | 15/04/2017
|  | 15 | RM approved label found in dust bin | Mr. Sharad / Mr. Malay | Change the practise, RM label to be crossed by marker / pen, do not remove from empty drums / bag for sending at scap yard. | 2017-12-04
|  | 16 | Green/ Red marking on magnehelic gauges | Mr. Malay | Remove / clean marking on magnehelic gauges | 14/04/2017
|  | 17 | unclean room - zebra tape demarketion for unclean utinsils / equipments | Mr. Malay | Fix zebra tape | 19/04/2017
|  | 18 | No sufficient space to palce the Utensils at right corner of clean bin rack | Mr. Malay / Mr. Jignesh | clean bin rack - Support and label to be shift to left corner | 14/04/2017
|  | 19 | No ID no. on lids of bins | Mr. Malay / Mr. Jignesh | to be fix the ID no. , ID no. should be same for indudial bin and its lid. | 14/04/2017
|  | 20 | temp indicator of ganson mixer is not available at place | Mr. Sharad / Mr. Jignesh / Mr. Sunil | temp indicator to be fix | 17/04/2017
|  | 21 | Genson - NOT IN USE label not have sign and date | Mr. Sharad / Mr. Malay | Genson - NOT IN USE label fix with sign and date | 17/04/2017
|  | 22 | inspection light box D2 wiring to be neat, lux light & size check | Mr. Sharad / Mr. Jignesh / Mr. Sunil | inspection light box of D1 & D2 make with same size with proper colour & lux light | 19/04/2017
|  | 23 | Nerrow down the weight limit of F.G.shipper | Mr. Sharad / Mr. Malay | work out for making nerrow range ofF.G. shipper weight | 15/04/2017
|  | 24 | level of weighing balance not proper | Mr. Sharad / Mr. Malay | level of all balance to be checked with spirit level | 2017-12-04
|  | 25 | location of humidity meter is not as per hotest tepm. Mapping point | Mr. Sharad / Mr. Jignesh | re-locate the humidity meter at hottest position | 14/04/2017
|  | 26 | BMR - In BOM required quantity of purified water is Q.S. to 15000L but actual used quantity not mentioned. | Mr. Sharad | Mention the actual quantity in step wise instruction coloumn. | 20/04/2017
|  | 27 | BMR - Actual mixing time of RS not match with standered mixing time. | Mr. Sharad | Mixing time will be as per standered time. If any deviation happen then mark * and mention reason with sign & date. | 20/04/2017
|  | 28 | BMR - Equipment ID of DAL filling machine not mention | Mr. Sharad | Training to be given to write equipment no. in BMR | 20/04/2017
|  | 29 | Validation of DAL 60ml filling & packing - speed not mentioned, in place of speed sign available. | Mr. Sharad & Ms.Tejal | under discussion | 29/04/2017
|  | 30 | Validation of DAL 5L filling & packing - 02 page found empty, document not controlled. | Mr. Sharad & Ms.Tejal | under discussion | 29/04/2017
|  | 31 | CCFFW - Mixing time having range (10 to 30 min.) in validation no conclusion mentioned for mixing time | Mr. Sharad / Mr. Indresh | under discussion | 29/04/2017
|  | 32 | Filter C&S record - data not visible | Mr. Sharad | If any forign material fall on document & data is not visible then mark * and mention reason with sign & date | 22/04/2017
| Warehouse |  | Rodiant box no. 18 need proper grounding or change place | HR | Rodiant Box must be on surface need to set properly | 10/05/2017
|  | 1 | Pest control / Rodiant Box map Not display | HR | Pest control / Rodiant Box map put in display | 10/05/2017
|  | 2 | There is no sign board for man entry in warehouse | Stores | Put sign board for man entry | 05/05/2017
|  | 3 | Procedure for warehouse access not display | Stores | Gawning procedure for workman/staff/visitor need to display | 05/05/2017
|  | 4 | De-dusting area not identify | Stores | Designate one area for De-dusting | 30/04/2017
|  | 5 | Air Curtain not set Properly | Stores / Maintanance | Air curtain need to redress if needed replace | 10/05/2017
|  | 6 | Hoist area not clean | Stores | cleaning require in hoist area | 22/04/2017
|  | 7 | there is no sign board for undertest and approve material | Stores | Need to put sign board for undertest and approve material | 30/04/2017
|  | 8 | In packaging material's undertest label not showing Mfg & Exp date | Stores / IT | Need to correct undertest label in SAP | 30/04/2017
|  | 9 | Tublight cover very durty | Maintanance | Need to replace all tubelight cover | 30/04/2017
|  | 10 | Area Location chart not display | HR | Put Location Chart in display | 05/05/2017
|  | 11 | 2nd floor
|  | 12 | Trial Material not in one place | stores | Need to define area for trial material | 30/04/2017
|  | 13 | Rejection / Destruction SOP not discribe properly | stores | Rejection / destruction sop need to be revise | 05/05/2017
|  | 14 | Material issue practice is not as per sop | stores | Material Issue SOP need to revise | 05/05/2017
|  | 15 | In Receipt checklist Annexure-II there is no identity of invoice no. unable to trace perface receipt. | stores | Receipt check list and inward check list need to redefine | 10/05/2017