1) Calibration of WB-01 & WB-02 are completed on quarterly (April-2016,july-2016,Oct-2016,Jan-2017,march-2017) 2) Opportunity base PM of WB-01 is completed on dated 15 April 2017 in that include painting ,mud cleaning ,load cell cleaning & calibration done 3) WB- 02 is not showing operational confidant from date of commission ,same load cell , base plate leveling & calibration done dated on 25 Dec 2016 .Now is show operational confidant with respective to WBI-01 4) PM weighing scale completed on quarterly (April-2016,july-2016,Oct-2016,Jan-2017,march-2017) 5) Stamping & sealing from 3rd party completed quarterly in front of government officer ( Weight & measure dept) 6) To increase lifespan & accuracy of standard 20Kg weight ,provide SS box in audit area ,painting & new pallet 7) 20kg standard weight comparison done with M1 class weight ,same difference noted . 8) syndent tank weighing calibration done as per process dept requirement